## Readme
Background information and Instructions
LOQ definition and harmonised way of determination | Limit of quantification. Lowest level that has been tested during (initial) validation and meets the criteria for recovery, repeatability and identification see: https://ec.europa.eu/food/system/files/2020-01/pesticides_mrl_guidelines_wrkdoc_2019-12682.pdf and: https://ec.europa.eu/food/system/files/2021-02/pesticides_mrl_guidelines_2020-12830.pdf
LOD definition & harmonised way of estimation | Limit of detection. Estimated lowest level at which the analyte can be detected and in principle also identified (so: in MS/MS both quantifier and qualifier present and meeting ion ratio criteria). This is a method/technical limit. Note that the LOD for many matrices may vary from sample to sample and also depend on the condition of the instrument at the time of measurement. It is therefore in fact not a fixed exact value, but an indication. There are multiple approaches and ways of estimating the LOD. For practical reasons the S/N approach is used. The LOD is the level at which S/N for both quantifier and qualifier are at least 3. Note that S/N may depend on acquisition rate and smoothing of the data. Software-based S/N often yield rather high values = optimistic LODs. Be realistic and provide the LOD that can gnerally be achieved in the routine situation / during samples analysis
Identification | For identification SANTE/12682/2019 is followed. At least 2 transitions are measured. Identification is based on matching retention time and signals for both transitions. The retention time should be within ±0.1 min from the reference standard. The analyte peaks in the extracted ion chromatograms must overlap. The ion ratio of the transitions should be within ±30% of that of the reference standard(s). For HRMS, see SANTE document. Not complying with the ID criteria does not automatically disqualify the result, but the result should be flagged (see below)
Reporting results | Be critical on potential false positives due to lab background or instrument carry-over Only report when the negative control (e.g. reagent blank, blank sample) is <LOD. If the negative control is not <LOD, reporting is still possible when the following requirements are met: - the level in the negative control is consistent/stable - the concentration in the sample is at least 3x higher than the negative control. In this case, the level in the negative control needs to be subtracted from the sample result.
Entry in pesticide column |  | Empty cell in results table means: result not yet available (analysis not yet performed).
| n.a. | For samples or pesticides that can not be analysed (sample not received, not sufficient material, pesticide not included in method, etc): report as "n.a." (pesticide not analysed).
| - | means <LOD, i.e. the sample has been analysed for the pesticide, but it was not present in the sample/not detected in the sample.
| [number] | Pesticide present in the sample (>LOD or >LOQ) at the indicated concentration. Provide the concentration without rounding, as the data are used by WP2 for further calculations. In the "format number" function of Excel, you can restrict the number of digits displayed. (note: for reporting in deliverables/publications, 3 significant figures are used for values <1000, and whole number for values ≥1000 (e.g. 1.24, 12.4, 124, 1245)) Concentrations between LOD and LOQ are concidered indicative only, and are likely to have higher/unknown uncertainty. They will be flagged as "<LOQ" in the flag column (see below)
| < [number] | LOQ cannot be achieved. In case the established LOQ for a certain pesticide in a certain sample cannot be achieved, e.g. due to a strong interference, extreme ion suppression, too small amount of sample material, report as: "<x.x" where x.x is the estimated level that would have been detectable. Flag in the flag column (see below).
Flag column | The flag column after each pesticide is used to indicate deviations, abnormalities and remarks at an individual sample/pesticide level. The most common issues are provided in the drop down menu. If you want to provide additional remarks/clarification, please right-click in the flag cell, and use the excel "new note" function to add free text as note.
| <LOQ | use this flag when the concentration is below the method LOQ.
| ID: qualifier ND/INT | identification issue: the qualifier was not detected or strong interferences were observed for the qualifier transition
| ID: IR>30% | identification issue: the ion ratio criterion was not met.
| BG subtracted | the pesticide is present in the negative control. Result reported is the value after subtraction of the background.
| LOQ issue | use this flag in case the established LOQ for a certain pesticide in a certain sample cannot be achieved, e.g. due to a strong interference, extreme ion suppression, too small amount of sample material.
| other remark | Use this flag for any other abnormalities or deviations then the ones indicated above. Use the excel "new note" function to add free text as note
Reporting for QC sheet | Method | For each pesticide indicate the method used for analysis. We distinguish between the different methods by the instrumental analysis. At a sample matrix level, this should be unique enough (if not inform Anne). For pesticides not analysed, select "n.a."
| LOD | specify the LOD, taking into account the remarks above
| LOQ | specify the LOQ, as established during validation (see above)
| initial validation | provide here the outcome of the initial validation: average recovery, repeatability, and N = the number replicates used for these. Do this for each level included in the validation.
| on-going validation | provide here the outcome of the on-going validation, based on spiked samples included in each batch of sample analysis: average recovery, within-lab reproducibility, and the number of QCs samples used for these. Do this for each level included in the batch QC.
| remarks | provide any remark considered relevant for the quality and interpretation of the analysis results

## Indoor_dust
Indoor dust | Method | LOD | LOQ | µg/kg -1 | Rec (ave) | RSDr | N | µg/kg - 2 | Rec (ave) | RSDr | N | µg/kg - 3 | Rec (ave) | RSDr | N | µg/kg 1b | Rec (ave) | RSDwR | N | µg/kg 2b | Rec (ave) | RSDwR | N
|  |  |  | QC initial validation | [11 empty] | On-going validation (from batch QCs concurrently analysed with the samples)
|  | µg/kg | µg/kg | level-1 (column E; µg/kg) |  |  |  | level-2 (column I; µg/kg) |  |  |  | level-3 (column M; µg/kg) |  |  |  | level-1 (column Q; µg/kg) |  |  |  | level-2 (column U; µg/kg)
| Method | LOD | LOQ | µg/kg | Rec (ave) | RSDr | N | µg/kg | Rec (ave) | RSDr | N | µg/kg | Rec (ave) | RSDr | N | µg/kg | Rec (ave) | RSDwR | N | µg/kg | Rec (ave) | RSDwR | N
DDE p,p' | 7 GC-HRMS | 0.25658 | 1 | 1 | 86.5 | 6.4829 | 5 | 10 | 91.84 | 4.837 | 5 | 50 | 90.02 | 7.4193 | 5 | 1 | 79.1333 | 6.2583 | 5 | - | - | - | -
Glyphosate | 4 Gly/ampa | 20 | 50 | 10 | - | - | 5 | 50 | 84.5315 | 11.9048 | 5 | 100 | 98.8445 | 15.502 | 5 | 50 | 89.6667 | 14.9902 | 5 | - | - | - | -
Piperonyl butoxide | 1 LC-MS/MS pos | 0.11599 | 10 | 1 | - | - | 5 | 10 | 100 | 1.0851 | 5 | 50 | 103.2919 | 9.1218 | 4 | 10 | 107.5861 | 4.6244 | 5 | 50 | 76.3733 | 3.9004 | 4
Tebuconazole | 1 LC-MS/MS pos | 0.33308 | 10 | 1 | - | - | 5 | 10 | 90.4767 | 5.169 | 5 | 50 | 94.1562 | 12.6515 | 4 | 10 | 94.0332 | 7.9475 | 5 | 50 | 93.0674 | 9.5517 | 4
Azoxystrobin | 1 LC-MS/MS pos | 0.1532 | 10 | 1 | - | - | 5 | 10 | 91.0828 | 1.1999 | 5 | 50 | 100.9946 | 7.1245 | 4 | 10 | 109.5706 | 5.0887 | 5 | 50 | 91.967 | 5.1408 | 4
Chlorpyrifos/-methyl: TCPy | 2 LC-MS/MS neg | 0.077709 | 10 | 1 | - | - | 5 | 10 | 88.6138 | 9.3178 | 5 | 50 | 79.8102 | 5.4733 | 4 | 10 | 77.9018 | 4.1882 | 5 | 50 | 100 | 2.5378 | 4
Fipronil sulfone | 2 LC-MS/MS neg | 0.085245 | 10 | 1 | - | - | 5 | 10 | 70.2094 | 1.1012 | 5 | 50 | 97.8208 | 15.0314 | 4 | 10 | 109.9055 | 8.183 | 5 | 50 | 79.6797 | 5.408 | 4
AMPA | 4 Gly/ampa | 38 | 100 | 10 | - | - | 5 | 50 | 74.9079 | 52.3023 | 5 | 100 | 92.0569 | 6.077 | 5 | 100 | 88.8333 | 8.8184 | 5 | - | - | - | -
Hexachlorobenzene | 7 GC-HRMS | 0.011525 | 1 | 1 | 78.7784 | 10.4658 | 5 | 10 | 82.55 | 1.2589 | 5 | 50 | 76.825 | 6.6179 | 5 | 1 | 74 | 4.187 | 5 | - | - | - | -
Boscalid | 1 LC-MS/MS pos | 0.72448 | 10 | 1 | - | - | 5 | 10 | 100.169 | 7.919 | 5 | 50 | 97.9617 | 6.2851 | 4 | 10 | 96.0843 | 6.3273 | 5 | 50 | 91.0115 | 7.1589 | 4
Deltamethrin | 3 GC-MS/MS | 1.0449 | 10 | 1 | - | - | 5 | 10 | 88.6678 | 9.491 | 5 | 50 | 106.6344 | 8.164 | 4 | 10 | 82.3097 | 11.6191 | 5 | - | - | - | -
Fludioxonil | 2 LC-MS/MS neg | 0.069563 | 10 | 1 | - | - | 5 | 10 | 107.1381 | 1.4852 | 5 | 50 | 116.7267 | 2.6625 | 4 | 10 | 104.9804 | 7.071 | 5 | 50 | 100 | 3.4969 | 4
Fipronil | 2 LC-MS/MS neg | 0.12314 | 10 | 1 | - | - | 5 | 10 | 93.3498 | 0.95087 | 5 | 50 | 94.6913 | 11.7387 | 4 | 10 | 86.0624 | 7.9216 | 5 | 50 | 100 | 5.3002 | 4
2,4-D (free) | 2 LC-MS/MS neg | 3.3842 | 10 | 1 | - | - | 5 | 10 | 71.1018 | 8.5692 | 5 | 50 | 90.7656 | 17.7827 | 4 | 10 | 86.8481 | 8.7397 | 5 | 50 | 104.5038 | 2.5088 | 4
Propiconazole | 1 LC-MS/MS pos | 0.26805 | 10 | 1 | - | - | 5 | 10 | 98.4238 | 12.5135 | 5 | 50 | 94.5011 | 10.9188 | 4 | 10 | 90.3328 | 7.9042 | 5 | 50 | 96.9694 | 9.1783 | 4
Pirimiphos-methyl | 1 LC-MS/MS pos | 0.39569 | 1 | 1 | 105.6979 | 4.5006 | 5 | 10 | 93.0048 | 8.5361 | 5 | 50 | 89.8841 | 12.0108 | 4 | 10 | 91.2423 | 6.9446 | 5 | 50 | 90.8896 | 8.0672 | 4
Cypermethrin | 3 GC-MS/MS | 3.2457 | 10 | 1 | - | - | 5 | 10 | 97.8076 | 6.7309 | 5 | 50 | 98.1077 | 6.0373 | 4 | 10 | 87.8807 | 14.6443 | 5 | - | - | - | -
Metalaxyl (M) | 1 LC-MS/MS pos | 0.32263 | 1 | 1 | 100.8431 | 12.1398 | 5 | 10 | 103.8492 | 4.6064 | 5 | 50 | 91.155 | 10.1254 | 4 | 10 | 99.6632 | 3.9789 | 5 | 50 | 103.1735 | 3.6843 | 4
Propoxur | 1 LC-MS/MS pos | 0.17167 | 1 | 1 | 101.3612 | 7.6435 | 5 | 10 | 89.6767 | 11.1098 | 5 | 50 | 104.5469 | 9.3601 | 4 | 10 | 99.5929 | 5.6946 | 5 | 50 | 79.2064 | 19.2972 | 4
Chlorpyrifos | 3 GC-MS/MS | 0.3 | 1.2 | 1.2 | 110.2379 | 4.7117 | 5 | 12 | 104.8818 | 16.4244 | 5 | 60 | 104.0838 | 18.0689 | 4 | 12 | 95.9052 | 16.9808 | 5 | - | - | - | -
RSDr = Relative standard deviation intraday repeatability determined at low, interermediate, and high concentrations; RSDwr = relative standard deviation within laboratory reproducibility based on N field samples; average recovery (%); "_" not analysed for this matrix; L, low level; Int, intermediate level

## Crop
Crop |  |  |  | QC initial validation | [8 empty] | On-going validation (from batch QCs concurrently analysed with the samples)
|  | µg/kg dry weight | µg/kg dry weight | level-1 2 µg/kg dry weight |  |  | level-2 20 µg/kg dry weight |  |  | level-3 x.x µg/kg dry weight |  |  | level-1 x.x µg/kg dry weight |  |  | level-2 x.x µg/kg dry weight
| Method | LOD | LOQ | Rec (ave) | RSDr | N | Rec (ave) | RSDr | N | Rec (ave) | RSDr | N | Rec (ave) | RSDwR | N | Rec (ave) | RSDwR | N | Remarks
DDE p,p' | 3 GC-MS/MS | 0.33333 | 1 | 71 | 9 | 5 | 90 | 0 | 5 |  |  | 5
Glyphosate | 4 Gly/ampa | 16.6667 | 50 | 106 | 15 | 5 | 98 | 2 | 5 |  |  | 5
Piperonyl butoxide | 1 LC-MS/MS pos | 0.66667 | 2 | 98.9687 | 1.457 | 5 | 90.4011 | 0.90351 | 5 |  |  | 5
Tebuconazole | 1 LC-MS/MS pos | 0.66667 | 2 | 90.7541 | 12.7723 | 5 | 96.9041 | 0.776 | 5 |  |  | 5
Azoxystrobin | 1 LC-MS/MS pos | 0.66667 | 2 | 105.6864 | 10.1554 | 5 | 93.1811 | 3.0583 | 5 |  |  | 5
Chlorpyrifos/-methyl: TCPy | 2 LC-MS/MS neg | 3.3333 | 10 | _ | _ | 5 | 85 | 5 | 5 |  |  | 5
Fipronil sulfone | 2 LC-MS/MS neg | 0.66667 | 2 | 98.6586 | 8.146 | 5 | 87 | 9.0618 | 5 |  |  | 5
AMPA | 4 Gly/ampa | 16.6667 | 50 | 99 | 18 | 5 | 95 | 7 | 5 |  |  | 5
Hexachlorobenzene | 3 GC-MS/MS | 0.33333 | 1 | 96.625 | 9.5587 | 5 | 127.7718 | 7.6552 | 5 |  |  | 5
Boscalid | 1 LC-MS/MS pos | 0.66667 | 2 | 91.5095 | 29.338 | 5 | 96.5722 | 9.5435 | 5 |  |  | 5
Deltamethrin | 3 GC-MS/MS | 0.33333 | 1 | 85.875 | 12.6383 | 5 | 107.6708 | 10.4311 | 5 |  |  | 5
Fludioxonil | 2 LC-MS/MS neg | 0.66667 | 2 | 98.2515 | 14 | 5 | 99 | 17.844 | 5 |  |  | 5 | [6 empty] | for oil seed and olives LOQ is 10
Fipronil | 2 LC-MS/MS neg | 0.66667 | 2 | 135.9497 | 36 | 5 | 105.9497 | 11.4236 | 5 |  |  | 5
2,4-D (free) | 2 LC-MS/MS neg | 0.66667 | 2 | 132.8153 | 24 | 5 | 102.8153 | 3.2388 | 5 |  |  | 5
Propiconazole | 1 LC-MS/MS pos | 0.66667 | 2 | 88.2679 | 3.5548 | 5 | 83.7577 | 0.38593 | 5 |  |  | 5
Pirimiphos-methyl | 1 LC-MS/MS pos | 0.66667 | 2 | 79.5263 | 4.2433 | 5 | 84.9778 | 4.2699 | 5 |  |  | 5
Cypermethrin | 3 GC-MS/MS | 0.33333 | 1 | 77.375 | 14.3003 | 5 | 88.0259 | 13.9261 | 5 |  |  | 5 | [6 empty] | for oil seed and olives LOQ is 10
Metalaxyl (M) | 1 LC-MS/MS pos | 0.66667 | 2 | 103.7769 | 3.1385 | 5 | 96.4066 | 0.69087 | 5 |  |  | 5
Propoxur | 1 LC-MS/MS pos | 0.66667 | 2 | 88.2679 | 3.5548 | 5 | 83.7577 | 0.38593 | 5 |  |  | 5
Chlorpyrifos | 3 GC-MS/MS | 0.33333 | 1 | 77.375 | 14.3003 | 5 | 88.0259 | 13.9261 | 5 |  |  | 5 | [6 empty] | for oil seed and olives LOQ is 10
Note on E5: Zuzana Bílková: level 1 for GLY/AMPA 50 ug/kg
Note on H5: Zuzana Bílková: level 2 for GLY/AMPA 500 ug/kg
Note on E11: Zuzana Bílková: level 1 for GLY/AMPA 50 ug/kg
Note on H11: Zuzana Bílková: level 2 for GLY/AMPA 500 ug/kg

## Earthworm
Earthworms |  |  |  | QC initial validation | [8 empty] | On-going validation (from batch QCs concurrently analysed with the samples)
|  | µg/kg dry weight | µg/kg dry weight | level-1 20 µg/kg dry weight |  |  | level-2 40 µg/kg dry weight |  |  | level-3 x.x µg/kg dry weight |  |  | level-1 x.x µg/kg dry weight |  |  | level-2 x.x µg/kg dry weight
| Method | LOD | LOQ | Rec (ave) | RSDr | N | Rec (ave) | RSDr | N | Rec (ave) | RSDr | N | Rec (ave) | RSDwR | N | Rec (ave) | RSDwR | N | Remarks
DDE p,p' | 3 GC-MS/MS | 3.3333 | 10 | 90 | 0 | 5 | 71 | 9 | 5 |  |  | 5
Glyphosate | 4 Gly/ampa | 66.6667 | 200 | 85 | 7 | 5 | 94 | 9 | 5 |  |  | 5
Piperonyl butoxide | 1 LC-MS/MS pos | 6.6667 | 20 | 71.7986 | 5.0166 | 5 | 68.3151 | 0.65449 | 5 |  |  | 5
Tebuconazole | 1 LC-MS/MS pos | 6.6667 | 20 | 71.7986 | 5.0166 | 5 | 68.3151 | 0.65449 | 5 |  |  | 5
Azoxystrobin | 1 LC-MS/MS pos | 6.6667 | 20 | 87.6472 | 19.9716 | 5 | 85.7991 | 1.2969 | 5 |  |  | 5
Chlorpyrifos/-methyl: TCPy | 2 LC-MS/MS neg | 13.3333 | 40 | _ | _ | 5 | 75.3791 | 20.2454 | 5 |  |  | 5
Fipronil sulfone | 2 LC-MS/MS neg | 6.6667 | 20 | 83.0709 | 1.3836 | 5 | 73.9016 | 15.2089 | 5 |  |  | 5
AMPA | 4 Gly/ampa | 66.6667 | 200 | 79 | 10 | 5 | 88 | 15 | 5 |  |  | 5
Hexachlorobenzene | 3 GC-MS/MS | 3.3333 | 10 | 88 | 18.7794 | 5 | 80.2333 | 0.91773 | 5 |  |  | 5
Boscalid | 1 LC-MS/MS pos | 6.6667 | 20 | 81.0509 | 39.184 | 5 | 75.7911 | 3.3194 | 5 |  |  | 5
Deltamethrin | 3 GC-MS/MS | 6.6667 | 20 | _ | _ | 5 | 74 | 12.3645 | 5 |  |  | 5
Fludioxonil | 2 LC-MS/MS neg | 13.3333 | 40 | _ | _ | 5 | 64.2487 | 12 | 5 |  |  | 5
Fipronil | 2 LC-MS/MS neg | 6.6667 | 20 | 83.5231 | 2.0152 | 5 | 77.2935 | 7.1811 | 5 |  |  | 5
2,4-D (free) | 2 LC-MS/MS neg | 6.6667 | 20 | 74.9523 | 6.7305 | 5 | 77.5389 | 11.952 | 5 |  |  | 5
Propiconazole | 1 LC-MS/MS pos | 6.6667 | 20 | 84.7681 | 11.1355 | 5 | 83.6602 | 1.7257 | 5 |  |  | 5
Pirimiphos-methyl | 1 LC-MS/MS pos | 6.6667 | 20 | 86.7287 | 9.4817 | 5 | 82.8419 | 1.8358 | 5 |  |  | 5
Cypermethrin | 3 GC-MS/MS | 6.6667 | 20 | _ | _ | 5 | 109.6 | 10.5122 | 5 |  |  | 5
Metalaxyl (M) | 1 LC-MS/MS pos | 6.6667 | 20 | 88.0501 | 2.8265 | 5 | 83.0629 | 0.48079 | 5 |  |  | 5
Propoxur | 1 LC-MS/MS pos | 6.6667 | 20 | 78.1301 | 6.0082 | 5 | 72.2121 | 2.1538 | 5 |  |  | 5
Chlorpyrifos | 3 GC-MS/MS | 3.3333 | 10 | 84 | 12.083 | 5 | 76 | 14.547 | 5 |  |  | 5
Note on E5: Zuzana Bílková: level 1 for GLY/AMPA 200 ug/kg
Note on H5: Zuzana Bílková: level 2 for GLY/AMPA 2000 ug/kg
Note on E11: Zuzana Bílková: level 1 for GLY/AMPA 200 ug/kg
Note on H11: Zuzana Bílková: level 2 for GLY/AMPA 2000 ug/kg

## Soil
Soil |  |  |  | QC initial validation | [8 empty] | On-going validation (from batch QCs concurrently analysed with the samples)
|  | µg/kg dry weight | µg/kg dry weight | level-1 4 µg/kg dry weight |  |  | level-2 40 µg/kg dry weight |  |  | level-3 x.x µg/kg dry weight |  |  | level-1 x.x µg/kg dry weight |  |  | level-2 x.x µg/kg dry weight
| Method | LOD | LOQ | Rec (ave) | RSDr | N | Rec (ave) | RSDr | N | Rec (ave) | RSDr | N | Rec (ave) | RSDwR | N | Rec (ave) | RSDwR | N | Remarks
DDE p,p' | 3 GC-MS/MS | 0.33333 | 1 | 90 | 0 | 5 | 71 | 9 | 5 |  |  | 5
Glyphosate | 4 Gly/ampa | 16.6667 | 50 | 85 | 7 | 5 | 94 | 9 | 5 |  |  | 5
Piperonyl butoxide | 1 LC-MS/MS pos | 1.3333 | 4 | 84.9373 | 8 | 5 | 83.3486 | 4.9282 | 5 |  |  | 5
Tebuconazole | 1 LC-MS/MS pos | 1.3333 | 4 | 71.013 | 8.116 | 5 | 83.9446 | 12 | 5 |  |  | 5
Azoxystrobin | 1 LC-MS/MS pos | 1.3333 | 4 | 91.1602 | 3.7704 | 5 | 81.5035 | 2.7706 | 5 |  |  | 5
Chlorpyrifos/-methyl: TCPy | 2 LC-MS/MS neg | 6.6667 | 20 | _ | _ | 5 | 85 | 5 | 5 |  |  | 5
Fipronil sulfone | 2 LC-MS/MS neg | 1.3333 | 4 | 98.6586 | 8.146 | 5 | 87 | 9.0618 | 5 |  |  | 5
AMPA | 4 Gly/ampa | 16.6667 | 50 | 79 | 10 | 5 | 88 | 15 | 5 |  |  | 5
Hexachlorobenzene | 3 GC-MS/MS | 0.33333 | 1 | 88 | 18.7794 | 5 | 80.2333 | 0.91773 | 5 |  |  | 5
Boscalid | 1 LC-MS/MS pos | 1.3333 | 4 | 83 | 15 | 5 | 93.2636 | 21.8418 | 5 |  |  | 5
Deltamethrin | 3 GC-MS/MS | 3.3333 | 10 | _ | _ | 5 | 74 | 12.3645 | 5 |  |  | 5
Fludioxonil | 2 LC-MS/MS neg | 1.3333 | 4 | 98.2515 | 14 | 5 | 99 | 17.844 | 5 |  |  | 5
Fipronil | 2 LC-MS/MS neg | 1.3333 | 4 | 135.9497 | 36 | 5 | 105.9497 | 11.4236 | 5 |  |  | 5
2,4-D (free) | 2 LC-MS/MS neg | 1.3333 | 4 | 132.8153 | 24 | 5 | 102.8153 | 3.2388 | 5 |  |  | 5
Propiconazole | 1 LC-MS/MS pos | 1.3333 | 4 | 69.4927 | 13.8865 | 5 | 71.467 | 10.3321 | 5 |  |  | 5
Pirimiphos-methyl | 1 LC-MS/MS pos | 1.3333 | 4 | 91.6377 | 5.3744 | 5 | 85.4407 | 1.0931 | 5 |  |  | 5
Cypermethrin | 3 GC-MS/MS | 3.3333 | 10 | _ | _ | 5 | 109.6 | 10.5122 | 5 |  |  | 5
Metalaxyl (M) | 1 LC-MS/MS pos | 1.3333 | 4 | 86.9163 | 5.7416 | 5 | 90.8625 | 2.0915 | 5 |  |  | 5
Propoxur | 1 LC-MS/MS pos | 1.3333 | 4 | 86.654 | 1.0183 | 5 | 76.9423 | 12.2862 | 5 |  |  | 5
Chlorpyrifos | 3 GC-MS/MS | 0.33333 | 1 | 84 | 12.083 | 5 | 76 | 14.547 | 5 |  |  | 5
Note on E5: Zuzana Bílková: level 1 for GLY/AMPA 50 ug/kg
Note on H5: Zuzana Bílková: level 2 for GLY/AMPA 500 ug/kg
Note on E11: Zuzana Bílková: level 1 for GLY/AMPA 50 ug/kg
Note on H11: Zuzana Bílková: level 2 for GLY/AMPA 500 ug/kg

## Sediment
Sediment |  |  |  | QC initial validation | [8 empty] | On-going validation (from batch QCs concurrently analysed with the samples)
|  | µg/kg dry weight | µg/kg dry weight | level-1 1.0 µg/kg dry weight |  |  | level-2 25 µg/kg dry weight |  |  | level-3 50 µg/kg dry weight |  |  | level-1 10 µg/kg dry weight |  |  | level-2 25µg/kg dry weight
| Method | LOD | LOQ | Rec (ave) | RSDr | N | Rec (ave) | RSDr | N | Rec (ave) | RSDr | N | Rec (ave) | RSDwR | N | Rec (ave) | RSDwR | N | Remarks
DDE p,p' | 3 GC-MS/MS | 0.090909 | 1 | 102 | 3.332% | 5 |  |  |  | 96 | 6.019% | 5 | 101 | 11% | 5 | 105 | 13% | 5
Glyphosate | 4 Gly/ampa | 3.62 | 10 | 87 | 14% | 9 | 90 | 6.999% | 9 | 93 | 7.381% | 9 | 100.8 | 13% | 2 | 95.39 | 4% | 2
Piperonyl butoxide | 1 LC-MS/MS pos | 0.115 | 1 | 91 | 3.026% | 5 | 74 | 17% | 5 | 85 | 6.07% | 5 | 89.8367 | 6.731% | 3 | 84.69 | 5.697% | 3
Tebuconazole | 1 LC-MS/MS pos | 3.574 | 5 | _ | _ | 5 | _ | _ | 5 | 103.33 | 10.08% | 5 |  |  | 3 | 110.57 | 13.62% | 3
Azoxystrobin | 1 LC-MS/MS pos | 0.06 | 1 | 107.13 | 1.412% | 5 | 84.004 | 17.13% | 5 | 96.498 | 7.674% | 5 | 107.8333 | 6.486% | 3 | 90.5633 | 3.353% | 3
Chlorpyrifos/-methyl: TCPy | 2 LC-MS/MS neg | 0.8 | 1 | 111.22 | 11.92% | 5 | 105.092 | 15% | 5 | 96.204 | 11.21% | 5 | 104.6033 | 14.36% | 3 | 109.22 | 14.08% | 3
Fipronil sulfone | 2 LC-MS/MS neg | 0.01 | 1 | 151.0675 | 5.514% | 4 | 155.24 | 4.315% | 3 | 152.634 | 3.643% | 5 | 132.127 | 28.37% | 3 | 121.6367 | 12.33% | 3 | 0.78813
AMPA | 4 Gly/ampa | 4.03 | 10 | 86 | 13.76% | 9 | 95 | 4.075% | 9 | 134 | 5.28% | 9 | 112.9 | 3% | 2 | 122.5 | 16% | 2
Hexachlorobenzene | 3 GC-MS/MS | 0.48 | 1 | 84 | 4.272% | 5 |  |  |  | 83 | 9.873% | 5 | 89.8 | 24% | 5 | 86 | 18% | 5
Boscalid | 1 LC-MS/MS pos | 0.52 | 1 | 111 | 3.934% | 5 | 83 | 12.03% | 5 | 97 | 8.01% | 5 | 99 | 12.1% | 3 | 94.7367 | 8.399% | 3
Deltamethrin | 3 GC-MS/MS | 0.72414 | 1 | 74 | 11.87% | 5 |  |  |  | 60 | 9.541% | 5 | 56 | 37% | 4 | 74 | 34% | 5 | 1.5385
Fludioxonil | 2 LC-MS/MS neg | 0.03 | 1 | 114.6825 | 10.58% | 4 | 111.4067 | 2.384% | 3 | 105.75 | 3.963% | 5 | 103.1233 | 14.82% | 3 | 97.1967 | 7.511%
Fipronil | 2 LC-MS/MS neg | 0.02 | 1 | 111.943 | 18.23% | 4 | 123.8658 | 11.05% | 3 | 111.345 | 8.063% | 5 | 121.0233 | 15.15% | 3 | 111.31 | 8.65% | 3
2,4-D (free) | 2 LC-MS/MS neg | 0.16 | 1 | 83.4075 | 11.03% | 4 | 100.51 | 2.384% | 3 | 95.822 | 6.702% | 5 | 89.9233 | 11.26% | 3 | 82.86 | 3.105% | 3 | * | correction for recovery if:
Propiconazole | 1 LC-MS/MS pos | 0.566 | 1 | 118 | 9.857% | 5 | 99 | 16% | 5 | 101 | 7.309% | 5 | 102.18 | 25.45% | 3 | 100.7033 | 15.48% | 3
Pirimiphos-methyl | 1 LC-MS/MS pos | 0.334 | 1 | 95 | 5.73% | 5 | 80 | 19% | 5 | 86 | 9% | 5 | 84.7767 | 4.822% | 3 | 86.5833 | 4.633% | 3
Cypermethrin | 3 GC-MS/MS | 0.325 | 1 | 97 | 17.66% | 5 |  |  |  | 74 | 8.484% | 5 | 101 | 35% | 5 | 81 | 25% | 5
Metalaxyl (M) | 1 LC-MS/MS pos | 0.121 | 1 | 101 | 4.295% | 5 | 81.488 | 14.53% | 5 | 92.752 | 9.271% | 5 | 93.54 | 2.309% | 3 | 89.99 | 3.611% | 3
Propoxur | 1 LC-MS/MS pos | 0.094 | 1 | 102 | 6% | 5 | 84 | 14.86% | 5 | 93 | 7.062% | 5 | 94.1367 | 12.35% | 3 | 92.9167 | 2.022% | 3
Chlorpyrifos | 3 GC-MS/MS | 0.085714 | 1 | 110 | 6.324% | 5 |  |  |  | 85 | 10.23% | 5 | 94 | 30% | 5 | 90 | 25% | 5
Note on K4: Osman, Rima: The values for recovery and RSD are for level 10 ug/kg
Note on K12: Osman, Rima: The values for recovery and RSD are for level 10 ug/kg
Note on K14: Osman, Rima: The values for recovery and RSD are for level 10 ug/kg
Note on K20: Osman, Rima: The values for recovery and RSD are for level 10 ug/kg
Note on K23: Osman, Rima: level 10 ug/kg is used here

## Water
Water |  |  |  | QC initial validation | [11 empty] | On-going validation (from batch QCs concurrently analysed with the samples)
|  | ng/L | ng/L | level-1 (column E; ng/L) |  |  |  | level-2 (column I; ng/L) |  |  |  | level-3 (column M; ng/L) |  |  |  | level-1 x.x ng/L |  |  | level-2 x.x ng/L
| Method | LOD | LOQ | ng/L | Rec (ave) | RSDr | N | ng/L | Rec (ave) | RSDr | N | ng/L | Rec (ave) | RSDr | N | Rec (ave) | RSDwR | N | Rec (ave) | RSDwR | N | Remarks
DDE p,p' | 7 GC-HRMS | 0.008054 | 0.5 | 0.5 | 88.6667 | 5.8098 | 3 | 10 | 109.6616 | 6.1981 | 5 | 100 | 107 | 8 | 5
Glyphosate | 4 Gly/ampa | 1.944 | 10 | 10 | 95 | 13 | 5 |  |  |  |  | 100 | 108 | 10 | 5
Piperonyl butoxide | 1 LC-MS/MS pos | 0.23095 | 5 | 5 | 91.1827 | 11.4792 | 5 | 25 | 76.213 | 17.7425 | 4 | 50 | 93.5606 | 19.9143 | 5
Tebuconazole | 1 LC-MS/MS pos | 0.41667 | 5 | 5 | 86.4345 | 11.9129 | 5 | 25 | 105.1169 | 7.8157 | 4 | 50 | 84.9307 | 13.0036 | 5
Azoxystrobin | 1 LC-MS/MS pos | 1.2581 | 5 | 5 | 91.4175 | 11.2386 | 5 | 25 | 97.8768 | 6.2342 | 4 | 50 | 96.0663 | 6.9148 | 5
Chlorpyrifos/-methyl: TCPy | 2 LC-MS/MS neg | 2.8191 | 5 | 5 | 93.5501 | 20.3798 | 5 | 25 | 117.1715 | 16.9046 | 4 | 50 | 70.671 | 18.6596 | 5
Fipronil sulfone | 2 LC-MS/MS neg | 0.084705 | 5 | 5 | 84.4139 | 17.27 | 5 | 25 | 111.1831 | 18.1281 | 4 | 50 | 96.898 | 16.6788 | 5
AMPA | 4 Gly/ampa | 1.1705 | 10 | 10 | 100 | 10 | 5 |  |  |  |  | 100 | 107 | 6 | 5
Hexachlorobenzene | 7 GC-HRMS | 0.00094586 | 0.5 | 0.5 | 87.2 | 2.7706 | 3 | 10 | 117.5928 | 1.7342 | 5 | 100 | 111.7376 | 6.7482 | 5
Boscalid | 1 LC-MS/MS pos | 1.6192 | 5 | 5 | 117.5693 | 18.1617 | 5 | 25 | 89.6227 | 12.45 | 4 | 50 | 96.7117 | 6.8584 | 5
Deltamethrin | 3 GC-MS/MS | 0.6 | 5 | 5 | 84.2 | 6.2507 | 5 | 25 | 89 | 13.6229 | 4 | 50 | 97.25 | 12.3785 | 5
Fludioxonil | 2 LC-MS/MS neg | 1.1585 | 5 | 5 | 100.6316 | 15.6951 | 5 | 25 | 92.2768 | 11.9283 | 4 | 50 | 137.8173 | 17.84 | 5
Fipronil | 2 LC-MS/MS neg | 0.26809 | 5 | 5 | 105.8192 | 14.3116 | 5 | 25 | 97.0984 | 18.1139 | 4 | 50 | 110.6994 | 6.185 | 5
2,4-D (free) | 2 LC-MS/MS neg | 1.2179 | 5 | 5 | 93.0312 | 18.0384 | 5 | 25 | 112.5649 | 8.3618 | 4 | 50 | 109.9852 | 17.6578 | 5
Propiconazole | 1 LC-MS/MS pos | 0.94012 | 5 | 5 | 108.0267 | 6.7177 | 5 | 25 | 93.9377 | 9.2167 | 4 | 50 | 94.5309 | 6.0557 | 5
Pirimiphos-methyl | 1 LC-MS/MS pos | 0.1675 | 5 | 5 | 100.0611 | 2.3474 | 5 | 25 | 94.439 | 7.6213 | 4 | 50 | 96.3122 | 5.0355 | 5
Cypermethrin | 3 GC-MS/MS | 0.83 | 5 | 5 | 93.45 | 11.4433 | 5 | 25 | 104.8125 | 7.7568 | 4 | 50 | 93.6 | 6.9614 | 5
Metalaxyl (M) | 1 LC-MS/MS pos | 0.38438 | 5 | 5 | 92.1869 | 7.6081 | 5 | 25 | 97.9394 | 5.9792 | 4 | 50 | 89.8219 | 8.1483 | 5
Propoxur | 1 LC-MS/MS pos | 0.95154 | 5 | 5 | 90.7073 | 8.6017 | 5 | 25 | 96.0659 | 7.1176 | 4 | 50 | 90.0761 | 4.7829 | 5
Chlorpyrifos | 3 GC-MS/MS | 0.65 | 6 | 6 | 92 | 10.1675 | 5 | 30 | 97.25 | 13.3379 | 4 | 60 | 91 | 13.9867 | 5

## A_Feed
Animal Feed |  |  |  | Gly/AMPA and GC: 10 ug/kg |  |  | Gly/AMPA: 20 ug/kg |  |  | Gly/AMPA and GC: 50 ug/kg |  |  | Gly/AMPA: 10 ug/kg |  |  | Gly/AMPA: 20 ug/kg
|  |  |  | QC initial validation | [8 empty] | On-going validation (from batch QCs concurrently analysed with the samples)
|  | µg/kg | µg/kg | level-1 1 µg/kg |  |  | level-2 5 µg/kg |  |  | level-3 10 µg/kg |  |  | level-1: 1 µg/kg |  |  | level-2: 5 µg/kg
| Method | LOD | LOQ | Rec (ave) | RSDr | N | Rec (ave) | RSDr | N | Rec (ave) | RSDr | N | Rec (ave) | RSDwR | N | Rec (ave) | RSDwR | N | Remarks
DDE p,p' | 3 GC-MS/MS | 2 | 10 | 112 | 18 | 3 |  |  |  | 115 | 14 | 3 | 98 | 12 | 5 | 83 | 12 | 5
Glyphosate | 4 Gly/ampa | 5 | 10 | 110 | 6.1 | 5 | 100 | 7.1 | 5 | 97 | 2.2 | 5 | 77.865 | 0.099893 | 2 | 92.805 | 9.0746 | 2
Piperonyl butoxide | 1 LC-MS/MS pos | 0.5 | 1 | 101.156 | 12.1445 | 5 | 97.224 | 4.9828 | 5 | 93.112 | 9.9069 | 5 | 91.3567 | 11.0047 | 4 | 100.2178 | 10.5004 | 4
Tebuconazole | 1 LC-MS/MS pos | 2.5 | 5 | - |  |  | 92.786 | 24.1221 | 5 | 85.346 | 24.2054 | 5 | - |  |  | 78.0133 | 17.8226 | 4
Azoxystrobin | 1 LC-MS/MS pos | 0.5 | 1 | 87.288 | 9.1765 | 5 | 89.072 | 9.0957 | 5 | 95.55 | 8.2661 | 5 | 105.7328 | 8.665 | 4 | 98.0223 | 17.6128 | 4
Chlorpyrifos/-methyl: TCPy | 2 LC-MS/MS neg | 5 | 10 | - |  |  | - |  |  | 89 | 12 | 5 | - |  | 4 | 84.61 | 26.6652 | 4
Fipronil sulfone | 2 LC-MS/MS neg | 0.4 | 1 | 90.398 | 8.4283 | 5 | 106.808 | 3.2847 | 5 | 93.412 | 2.5024 | 5 | 112.205 | 10.9578 | 4 | 100.0595 | 13.0827 | 4
AMPA | 4 Gly/ampa | 10 | 20 | nd |  |  | 91 | 12 | 5 | 103 | 8.4 | 5 | - |  |  | 112.485 | 2.7219 | 2 | The samples used for QCs in feed were different in each serie (depending on the matrix measured). The results are the average of the recoveries considering the Qcs from all the series.
Hexachlorobenzene | 3 GC-MS/MS | 2 | 10 | 96 | 4 | 3 |  |  |  | 98 | 2 | 3 | 78 | 31 | 5 | 75 | 7 | 5
Boscalid | 1 LC-MS/MS pos | 2.5 | 5 | 132.02 | 34.035 | 5 | 116.458 | 20.3527 | 5 | 112.108 | 9.6062 | 5 | 100.3918 | 13.9285 | 4 | 96.9098 | 9.6119 | 4
Deltamethrin | 3 GC-MS/MS | 4 | 10 | 104 | 22 | 3 |  |  |  | 106 | 28 | 3 | 131 | 23 | 5 | 82 | 21 | 5
Fludioxonil | 2 LC-MS/MS neg | 2.5 | 5 | - |  |  | 123.7 | 6.1243 | 5 | 97.614 | 11.8549 | 5 | - |  |  | 76.5 | 2.7545 | 4
Fipronil | 2 LC-MS/MS neg | 0.4 | 1 | 65.574 | 9.2273 | 5 | 100.004 | 4.9234 | 5 | 102.854 | 3.615 | 5 | 109.0075 | 6.0208 | 4 | 98.97 | 17.9315 | 4
2,4-D (free) | 2 LC-MS/MS neg | 0.4 | 1 | 54.438 | 6.6751 | 5 | 52.964 | 4.1072 | 5 | 49.136 | 12.8196 | 5 | 85.6095 | 24.5108 | 4 | 72.9872 | 31.605 | 4
Propiconazole | 1 LC-MS/MS pos | 2.5 | 5 | - |  |  | 106.024 | 14.0868 | 5 | 95.162 | 7.2227 | 5 | 101.1 | 11.0929 | 4 | 97.7075 | 14.6842 | 4
Pirimiphos-methyl | 1 LC-MS/MS pos | 0.5 | 1 | 123.49 | 19.0676 | 5 | 95.096 | 6.5114 | 5 | 98.268 | 13.2137 | 5 | 96 | 16 | 4 | 97 | 14 | 4
Cypermethrin | 3 GC-MS/MS | 5 | 10 | 92 | 13 | 3 |  |  |  | 98 | 18 | 3 | 166 | 22 | 5 | 100 | 18 | 5
Metalaxyl (M) | 1 LC-MS/MS pos | 0.5 | 1 | 102.136 | 12.1805 | 5 | 90.928 | 9.8302 | 5 | 95.404 | 8.2725 | 5 | 109.0825 | 18.3262 | 4 | 99.735 | 20.8317 | 4
Propoxur | 1 LC-MS/MS pos | 2.5 | 5 | - |  |  | 96.912 | 21.4752 | 5 | 87.374 | 21.2 | 5 | 97.415 | 6.4285 | 4 | 106.235 | 11.6789 | 4
Chlorpyrifos | 3 GC-MS/MS | 2 | 10 | 104 | 16 | 3 |  |  |  | 112 | 7 | 3 | 136 | 21 | 5 | 106 | 6 | 5
Note on N8: Dias, Jonatan: No data at this level
Note on N10: Dias, Jonatan: No data at this level
Note on N16: Dias, Jonatan: No data at this level

## A_Feces
Animal Feces |  |  |  | QC initial validation | [8 empty] | On-going validation (from batch QCs concurrently analysed with the samples)
|  | µg/kg | µg/kg | level-1: 1 µg/kg |  |  | level-2: 2.5 µg/kg |  |  | level-3: 5 µg/kg |  |  | level-1: 1 µg/kg |  |  | level-2: 5 µg/kg
| Method | LOD | LOQ | Rec (ave) | RSDr | N | Rec (ave) | RSDr | N | Rec (ave) | RSDr | N | Rec (ave) | RSDwR | N | Rec (ave) | RSDwR | N
DDE p,p' | 1 LC-MS/MS pos | 1 | 1 | 91.1667 | 5.9845 | 6 | 95 | 2.491 | 6 | 97.8333 | 3.3849 | 6 | 115.056 | 9.4943 | 5 | 87.936 | 9.09 | 5
Glyphosate | 4 Gly/ampa | 25 | 30 | - |  |  | 85 | 3.1 | 5 | 98 | 2.7 | 5 | 178 | 6.5 | 4 | - | - | -
Piperonyl butoxide | 1 LC-MS/MS pos | 1 | 1 | 91.1667 | 5.9845 | 6 | 95 | 2.491 | 6 | 97.8333 | 3.3849 | 6 | 115.056 | 9.4943 | 5 | 87.936 | 9.09 | 5
Tebuconazole | 1 LC-MS/MS pos | 2 | 2.5 | nd |  |  | 96.8333 | 10.3768 | 6 | 100.3333 | 13.588 | 6 | - |  |  | 73.0975 | 25.5244 | 5
Azoxystrobin | 1 LC-MS/MS pos | 0.4 | 1 | 83 | 8.4 | 6 | 96 | 4.4 | 6 | 100 | 2.5 | 6 | 94 | 10.4732 | 5 | 88.126 | 4.1613 | 5
Chlorpyrifos/-methyl: TCPy | 2 LC-MS/MS neg | 0.5 | 1 | 83.1667 | 17.0626 | 6 | 88.5 | 15.7342 | 6 | 87.5 | 5.0467 | 6 | 81.5 | 21.7092 | 5 | 92.406 | 17.8872 | 5
Fipronil sulfone | 2 LC-MS/MS neg | 0.1 | 1 | 92.8157 | 4.4922 | 6 | 101.2262 | 8.72 | 6 | 104.2 | 9.9982 | 6 | 107.3989 | 6.8483 | 19 | 91.4668 | 8.5067 | 19
AMPA | 4 Gly/ampa | 10 | 25 | 125 | 9 | 5 | 109 | 8 | 5 | 105 | 2 | 5 | - |  |  | 95 | 15 | 14
Hexachlorobenzene | 3 GC-MS/MS | 0.15 | 1 | 71.3177 | 7.7157 | 5 | - | - | - | 63.9582 | 6.1165 | 5 | 77 | 16 | 29 | 66 | 14 | 34
Boscalid | 1 LC-MS/MS pos | 4 | 5 | nd |  | 6 | nd | - | 6 | 100.6188 | 7.1855 | 6 | - | - | - | 106.2781 | 24.3886 | 19
Deltamethrin | 3 GC-MS/MS | 0.4 | 1 | * | * | * | - | - | - | 79.7048 | 17.0572 | 5 | - | - | - | 104 | 30 | 34
Fludioxonil | 2 LC-MS/MS neg | 0.5 | 1 | 109.1949 | 9.892 | 6 | 103.7952 | 6.0937 | 6 | 98.8525 | 6.9507 | 6 | 107.67 | 28.3657 | 19 | 92.5609 | 24.0598 | 19
Fipronil | 2 LC-MS/MS neg | 0.1 | 1 | 98.6705 | 4.3312 | 6 | 98.1398 | 6.0045 | 6 | 99.3489 | 6.9022 | 6 | 96.07 | 7.6432 | 19 | 90.7374 | 9.7762 | 19
2,4-D (free) | 2 LC-MS/MS neg | 0.5 | 1 | 57.1615 | 8.487 | 6 | 55.6312 | 2.9499 | 6 | 57.8343 | 8.0742 | 6 | 63.1867 | 20.2991 | 19 | 54.8237 | 20.9358 | 19
Propiconazole | 1 LC-MS/MS pos | 0.5 | 1 | 90.3166 | 18.7573 | 6 | 83.8221 | 9.3945 | 6 | 92.179 | 6.5391 | 6 | 84.5281 | 25.7981 | 19 | 75.0024 | 33.4604 | 19
Pirimiphos-methyl | 1 LC-MS/MS pos | 0.5 | 1 | 86.0511 | 15.8646 | 6 | 93.7764 | 5.6405 | 6 | 90.4263 | 0.92685 | 6 | 95.4768 | 12.82 | 19 | 82.1558 | 14.8562 | 19
Cypermethrin | 3 GC-MS/MS | 0.5 | 1 | * | * | * |  |  |  | 70.8657 | 6.621 | 4 | - |  |  | 109 | 16 | 34
Metalaxyl (M) | 1 LC-MS/MS pos | 0.5 | 1 | 100.0605 | 6.2463 | 6 | 97.5443 | 1.6003 | 6 | 94.0302 | 5.5044 | 6 | 92.3553 | 10.936 | 19 | 86.7411 | 12.3822 | 19
Propoxur | 1 LC-MS/MS pos | 0.5 | 1 | 98.5812 | 9.2146 | 6 | 102.4053 | 7.379 | 6 | 98.984 | 6.5427 | 6 | 93.6989 | 17.1577 | 19 | 89.96 | 20.1641 | 19
Chlorpyrifos | 3 GC-MS/MS | 0.23636 | 1 | 85.5524 | 7.9847 | 5 | - | - | - | 93.1043 | 12.3642 | 4 | 116 | 24 | 27 | 108 | 57 | 29
Note on H2: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: For GC level 2 is 10 ug/kg
Note on Q2: Dias, Jonatan: For GC level 2 is 10 ug/kg
Note on E5: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The blank was contaminated with glyphosate (around 170 ug/kg), so the results for the validation are substracted from the blank. At 10 ug/kg the calculations were not good, so it was excluded.
Note on N5: Dias, Jonatan: This is not recovery, but the average concentration in ug/kg of the blank evaluated in each serie.
Note on N7: Dias, Jonatan: not validated at this level
Note on N13: Dias, Jonatan: not validated at this level
Note on C14: Mol, Hans: 1 transition (ID supported by 2 isomer peaks)
Note on N14: Dias, Jonatan: Present in the blank. Calculation not posible
Note on C20: Mol, Hans: 1 transition (ID supported by 2 isomer peaks)
Note on N20: Dias, Jonatan: Present in the blank. Calculation not posible

## A_Plasma
Animal Plasma |  |  |  | QC initial validation | [8 empty] | On-going validation (from batch QCs concurrently analysed with the samples)
|  | ng/ml | ng/ml | level-1 x.x ng/ml |  |  | level-2 x.x ng/ml |  |  | level-3 x.xng/ml |  |  | level-1 x.x ng/ml |  |  | level-2 x.x ng/ml
| Method | LOD | LOQ | Rec (ave) | RSDr | N | Rec (ave) | RSDr | N | Rec (ave) | RSDr | N | Rec (ave) | RSDwR | N | Rec (ave) | RSDwR | N | Remarks
DDE p,p' | 3 GC-MS/MS | 0.1 | 0.1 | [15 empty] | measured in Limoges
Glyphosate | 4 Gly/ampa | 0.05 | 0.1 | [15 empty] | measured in Limoges
Piperonyl butoxide | 1 LC-MS/MS pos | 0.001 | 0.125 | 104.2 | 3.9 | 6 | 98.4 | 4.7 | 6 | 97.9 | 8.9 | 4 |  |  |  | 80.9558 | 20.951 | 43 | L1=0,125, L2=1, L3=10 ng/ml
Tebuconazole | 1 LC-MS/MS pos | 0.145 | 0.25 | 107.3 | 19.5 | 6 | 96.1 | 18.4 | 6 | 95 | 4.9 | 4 |  |  |  | 100.4514 | 20.3832 | 43 | L1=0,25, L2=1, L3=10 ng/ml
Azoxystrobin | 1 LC-MS/MS pos | 0.002 | 0.125 | 102.7 | 4.7 | 6 | 105.7 | 4.7 | 6 | 98.2 | 4.4 | 4 |  |  |  | 83.475 | 16.504 | 43 | L1=0,125, L2=1, L3=10 ng/ml
Chlorpyrifos/-methyl: TCPy | 2 LC-MS/MS neg | qualitative only | 1 | n.a. | n.a. | 6 | 134.9 | 46.74 | 6 | 103.4 | 16.27 | 4 | [6 empty] | qualitative only
Fipronil sulfone | 2 LC-MS/MS neg | 0.0001 | 0.25 | 113.2 | 4.5 | 6 | 103.8 | 4.5 | _ | 99.6 | 1.74 | 4 |  |  |  | 114.0786 | 10.9996 | 42 | L1=0,25, L2=1, L3=10 ng/ml
AMPA | 4 Gly/ampa | 0.1 | 0.1 | [15 empty] | measured in Limoges
Hexachlorobenzene | 3 GC-MS/MS | 0.1 | 0.1 | [15 empty] | measured in Limoges
Boscalid | 1 LC-MS/MS pos | 0.039 | 0.125 | 101.8 | 17.3 | 6 | 97.1 | 6.3 | 6 | 93.1 | 4.6 | 4 |  |  |  | 92.186 | 17.925 | 43 | L1=0,125, L2=1, L3=10 ng/ml
Deltamethrin | 3 GC-MS/MS | 0.5 | 0.1 | [15 empty] | measured in Limoges
Fludioxonil | 2 LC-MS/MS neg | 0.028 | 0.25 | 113.4 | 5.9 | 6 | 105.1 | 3.64 | 6 | 91 | 5.4 | 3 |  |  |  | 99.793 | 11.1662 | 43 | L1=0,125, L2=1, L3=10 ng/ml
Fipronil | 2 LC-MS/MS neg | 0.0004 | 0.25 | 104.3 | 5.55 | 6 | 109 | 4.6 | 6 | 96.9 | 1.98 | 4 |  |  |  | 110.3395 | 14.8805 | 43 | L1=0,25, L2=1, L3=10 ng/ml
2,4-D (free) | 2 LC-MS/MS neg | qualitative only | 0.5 | n.a. | n.a. | 6 | 108.6667 | 21.8195 | 6 | 100.5 | 10.6708 | 4 | [6 empty] | qualitative only
Propiconazole | 1 LC-MS/MS pos | 0.036 | 0.25 | 80.4 | 17.72 | 6 | 99.7 | 7.5 | 6 | 98.9 | 2.5 | 4 |  |  |  | 91.3023 | 21.6831 | 43 | L1=0,25, L2=1, L3=10 ng/ml
Pirimiphos-methyl | 1 LC-MS/MS pos | 0.01 | 0.125 | 105.3 | 11.1 | 6 | 98.9 | 4.8 | 6 | 100.5 | 2.6 | 4 |  |  |  | 99.0953 | 12.2937 | 43 | L1=0,125, L2=1, L3=10 ng/ml
Cypermethrin | 3 GC-MS/MS | 0.1 | [16 empty] | measured in Limoges
Metalaxyl (M) | 1 LC-MS/MS pos | 0.008 | 0.125 | 94 | 6.7 | 6 | 105.2 | 4.6 | 6 | 96.8 | 5.5 | 4 |  |  |  | 84.9395 | 20.8765 | 43 | L1=0,125, L2=1, L3=10 ng/ml
Propoxur | 1 LC-MS/MS pos | 0.002 | 0.125 | 99.6 | 3.6 | 6 | 107.2 | 4.9 | 6 | 107.2 | 3.8 | 4 |  |  |  | 103.6023 | 10.5938 | 43 | L1=0,125, L2=1, L3=10 ng/ml
Chlorpyrifos | 3 GC-MS/MS | 0.5 | 0.5 | [15 empty] | measured in Limoges
Missing data is given below as it was analysed in another laboratory (Limoge, France)
Tabel S1. Quality control data for LC-MS/MS analysis of pesticides and their metabolites in blood.
Pesticide or metabolite | RSDr | RSDr | RSDr | RSDwr
| low (%) | inter (%) | high (%) | (%)
Piperonyl butoxide | 3.9 | 4.7 | 8.9 | 21
Tebuconazole | 19.5a | 18.4 | 4.9 | 20.4
Azoxystrobin | 4.7 | 4.7 | 4.4 | 16.5
Fipronil sulfone | 4.5 | 4.5 | 1.7 | 11.0b
Boscalid | 17.3 | 6.3 | 4.6 | 17.9
Fludioxonil | 5.9 | 3.6 | 5.4 | 11.2
Fipronil | 5.6a | 4.6 | 2 | 14.9
Propiconazole | 17.7a | 7.5 | 2.5 | 21.7
Pirimiphos-methyl | 11.1 | 4.8 | 2.6 | 12.3
Metalaxyl (M) | 6.7 | 4.6 | 5.5 | 20.9
Propoxur | 3.6 | 4.9 | 3.8 | 10.6
RSDr = Relative standard deviation intraday repeatability determined at low = 0.125 mg/mL, inter = 1.0 ng/mL and high = 10.0 ng/mL concentrations; RSDwr = relative standard deviation within laboratory reproducibility based on 43 field samples; a low = 0.25 ng/mL; N = 42
Tabel S2. QC control data for GC-MS/MS analysis of pesticides and their metabolites in blood.
Pesticide or metabolite | RSDr low (%) | RSDr inter (%) | RSDr high (%) | RSDwr low (%) | RSDwr inter (%) | RSDwr high (%)
DDE p,p' | 4.6 | 7.9 | 1.2 | 12.6 | 5.2 | 5.6
Glyphosate | 10.2 | 5.3 | 4.7 | 10.9 | 13.8 | 8
Chlorpyrifos methyl TCPy | 1.9 | 1.6 | 1.4 | 3.8 | 2.8 | 3.4
AMPA | 11.7 | 8.5 | 10.4 | 20.4 | 7.9 | 20.1
Hexachlorobenzene | 8.3 | 2.2 | 0.5 | 10.2 | 2 | 1.8
Deltamethrin | 4.9 | 5.1 | 3 | 13.2 | 8.5 | 5.4
2,4-D | 9.8 | 10.2 | 3.4 | 24.5 | 10.2 | 5.3
Cypermethrin | 6.8 | 5.4 | 2.7 | 10.9 | 8.9 | 7.9
Chlorpyrifos | 2.4 | 0.6 | 1.1 | 3.3 | 2.1 | 1.9
RSDr = Relative standard deviation intraday repeatability; RSDwr = relative standard deviation within laboratory reproducibility; low = lower limit of quantification (LLOQ), inter = intermediate concentration; high = higher limit of quantification (HLOQ)

## A_Urine
Animal Urine |  |  |  | QC initial validation | [8 empty] | On-going validation (from batch QCs concurrently analysed with the samples)
|  | ng/mL | ng/mL | level-1 0.1 ng/mL |  |  | level-2 0.5 ng/mL |  |  | level-3 2.5 ng/mL |  |  | level-1 0.1 ng/mL |  |  | level-2 0.5 ng/mL
| Method | LOD | LOQ | Rec (ave) | RSDr | N | Rec (ave) | RSDr | N | Rec (ave) | RSDr | N | Rec (ave) | RSDwR | N | Rec (ave) | RSDwR | N | Remarks
DDE p,p' |  | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
Glyphosate | 4 Gly/ampa |  | 0.1 | 100 | 14 | 5 | 94 | 44 | 5 | 97 | 2.9 | 5 | 100 | 22 | 27 | 103 | 16 | 27
Piperonyl butoxide |  | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
Tebuconazole-OH | 1 LC-MS/MS pos | 0.05 | 0.1 | 84 | 14.4 | 5 | 105 | 6.1 | 5 | 111 | 3.3 | 5 | 102 | 15.7 | 14 | 102 | 15.4 | 14
Azoxystrobin |  |  |  | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
Chlorpyrifos-methyl-desmethyl | 2 LC-MS/MS neg | 0.05 | 0.1 | 99 | 4 | 5 | 99 | 5 | 5 | 100 | 5 | 5 | 94 | 25.7 | 14 | 97 | 17 | 14
Fipronil sulfone | 2 LC-MS/MS neg | 0.05 | 0.1 | 118 | 2.9 | 5 | 111 | 2.5 | 5 | 117 | 1.9 | 5 | 112 | 15.8 | 14 | 99 | 17.1 | 14
AMPA | 4 Gly/ampa |  | 0.2 | 90 | 11 |  | 98 | 15 |  | - | - | - | 108 | 21 | 27 | 106 | 16 | 27
Hexachlorobenzene |  | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
Boscalid-OH | 1 LC-MS/MS pos | 0.05 | 0.1 | 103 | 9 | 5 | 113 | 10 | 5 | 114 | 3 | 5 | 87 | 18.4 | 14 | 90 | 20.4 | 14
Deltamethrin |  | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
Fludioxonil |  | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
Fipronil | 2 LC-MS/MS neg | 0.05 | 0.1 | 90 | 3 | 5 | 99 | 1.9 | 5 | 98 | 2.2 | 5 | 103 | 15.6 | 14 | 93 | 10.9 | 14
2,4-D (free) | 2 LC-MS/MS neg | 0.05 | 0.1 | 109 | 19 | 5 | 105 | 7 | 5 | 102 | 5 | 5 | 104 | 11.7 | 14 | 95 | 22.9 | 14
Propiconazole |  | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
Pirimiphos-methyl |  | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
Cypermethrin |  | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
Metalaxyl (M) |  | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
Propoxur |  | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
Chlorpyrifos |  | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
lambda-Cyhalothrin |  | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
Pyrimethanil-OH M605F002 | 1 LC-MS/MS pos | 0.05 | 0.1 | 90 | 22.9 | 5 | 103 | 9.1 | 5 | 92 | 3.3 | 5 | 109 | 10.8 | 14 | 98 | 10.3 | 14
Most pesticides are metabolized upon uptake. Specifically, for those detected in diet, WFSR checked the urinary biomarkers of exposure that may be detected. Among these substances, the lab team (Wageningen Food Safety Research, WFSR) checked if their analytical reference standard is (commercially) available.If available, the quantitative analysis of the substance is carried out. Due to these limitations, some substances were not analysed.
Note on F11: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Within-lab reproducibility
Note on I11: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Within-lab reproducibility

## A_Wristbands
Animal Wristband |  |  |  | QC initial validation | [8 empty] | On-going validation (from batch QCs concurrently analysed with the samples)
|  | µg/kg weight | µg/kg weight | level-1 1µg/kg weight |  |  | level-2 5.0 µg/kg weight |  |  | level-3 x.x µg/kg weight |  |  | level-1 1µg/kg weight |  |  | level-2 5.0 µg/kg weight
| Method | LOD | LOQ | Rec (ave) | RSDr | N | Rec (ave) | RSDr | N | Rec (ave) | RSDr | N | Rec (ave) | RSDwR | N | Rec (ave) | RSDwR | N | Remarks
DDE p,p' | GC-MS | 0.2 | - | [6 empty] | - | - | - | 114.765 | 58.02% | 8 | 87.5413 | 70.71% | 9
Glyphosate |  | - | - | [6 empty] | - | - | -
Piperonyl butoxide | 1 LC-MS/MS pos | 0.5 | 1 | 86.16 | 3.037% | 3 | 84.86 | 3.785% | 3 | - | - | - | 78.1346 | 14.93% | 13 | 76.9708 | 13.96% | 13
Tebuconazole | 1 LC-MS/MS pos | 2.5 | 5 | nd |  | 3 | 69.7867 | 31.94% | 3 | - | - | - | 75.442 | 40.38% | 13 | 73.3546 | 26.88% | 13
Azoxystrobin | 1 LC-MS/MS pos | 0.5 | 1 | 100.18 | 1.451% | 3 | 95.86 | 3.72% | 3 | - | - | - | 85.5292 | 15.17% | 12 | 83.025 | 13.02% | 12
Chlorpyrifos/-methyl: TCPy | 2 LC-MS/MS neg | 0.5 | 1 | 83.0433 | 6.213% | 3 | 82.4733 | 4.202% | 3 | - | - | - | 77.98 | 14.16% | 13 | 74.2154 | 16.58% | 13
Fipronil sulfone | 2 LC-MS/MS neg | 0.1 | 1 | 68.0333 | 2.403% | 3 | 63.3667 | 3.426% | 3 | - | - | - | 78.2546 | 12.04% | 13 | 76.3446 | 10.51% | 13
AMPA |  | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | recovery lower then 30% qual, recovery higher then 30% and RSD lower then 30% accepted , RSD higher then 30% increas the LOQ
Hexachlorobenzene | GC-MS |  | n.f.r | - | - | - | - | - | - | - | - | - | - | - | - | - | - | -
Boscalid | 1 LC-MS/MS pos | 0.2 | 1 | 102.1867 | 10.48% | 3 | 96.5133 | 2.91% | 3 | - | - | - | 82.7531 | 20.76% | 13 | 82.2454 | 20.11% | 13
Deltamethrin | GC-MS | 0.2 | - | [6 empty] | - | - | - | 139.545 | 54.56% | 8 | 106.3147 | 40.37% | 9
Fludioxonil | 2 LC-MS/MS neg | 0.2 | 1 | 60.78 | 5.07% | 3 | 61.8467 | 18.02% | 3 | - | - | - | 74.4815 | 13.76% | 13 | 73.4154 | 10.69% | 13
Fipronil | 2 LC-MS/MS neg | 0.1 | 1 | 71.77 | 1.738% | 3 | 69.9833 | 1.569% | 3 | - | - | - | 78.0623 | 13.74% | 13 | 78.9792 | 11.9% | 13
2,4-D (free) | 2 LC-MS/MS neg | 0.2 | 1 | 72.6733 | 9.282% | 3 | 90.73 | 13.92% | 3 | - | - | - | 44.4523 | 16.39% | 13 | 46.3246 | 9.692% | 13
Propiconazole | 1 LC-MS/MS pos | 0.5 | 1 | 100.1567 | 17.19% | 3 | 96.2733 | 8.247% | 3 | - | - | - | 79.6046 | 13.21% | 13 | 79.8254 | 15.81% | 13
Pirimiphos-methyl | 1 LC-MS/MS pos | 0.2 | 1 | 81.4433 | 5.377% | 3 | 82.89 | 4.018% | 3 | - | - | - | 74.8708 | 15.03% | 13 | 76.05 | 13.19% | 13
Cypermethrin | GC-MS | 0.2 | [7 empty] | - | - | - | 77.09 | 50.09% | 8 | 76.2418 | 47.99% | 9
Metalaxyl (M) | 1 LC-MS/MS pos | 0.2 | 1 | 96.4633 | 7.662% | 3 | 87 | 1.528% | 3 | - | - | - | 76.7146 | 14.33% | 13 | 76.3715 | 12.22% | 13
Propoxur | 1 LC-MS/MS pos | 0.2 | 1 | 90.22 | 5.49% | 3 | 92.0333 | 5.622% | 3 | - | - | - | 78.2131 | 20.88% | 13 | 74.0838 | 13.51% | 13
Chlorpyrifos | GC-MS | 0.2 | - | [6 empty] | - | - | - | 183.065 | 55.09% | 8 | 139.9164 | 55.52% | 9